## Slide 1: Laser Safety
The George Washington University Office of Laboratory Safety Environmental Health and Safety
Notes: <number>

## Slide 2: LASER
LASER stands for:
Light
Amplification by the
Stimulated
Emission of
Radiation
Notes: <number>

## Slide 3: Laser Light
Laser light
  is monochromatic, unlike ordinary light which is made of a spectrum of many wavelengths. Because the light is all of the same wavelength, the light waves are said to be synchronous.
  is directional and focused so that it does not spread out from the point of origin.
Synchronous, monochromatic, directional light waves
Asynchronous, multi-directional light.
Notes: <number>

## Slide 4: Uses of Lasers
Lasers are used in industry, communications, military, research and medical applications.
At university, lasers are used in both research and medical procedures.
Notes: <number>

## Slide 5: How a Laser Works
A laser consists of an optical cavity, a pumping system, and a lasing medium.
The optical cavity contains the media to be excited with mirrors to redirect the produced photons back along the same general path.
The pumping system uses various methods to raise the media to the lasing state.
The laser medium can be a solid (state), gas, liquid dye, or semiconductor.
Source: OSHA Technical Manual, Section III: Chapter 6, Laser Hazards.
Notes: <number>

## Slide 6: Types of Lasers
which part of the electromagnetic spectrum is represented:
  Infrared
  Visible Spectrum
  Ultraviolet
the length of time the beam is active:
  Continuous Wave
  Pulsed
  Ultra-short Pulsed
Lasers can be described by:
Notes: <number>

## Slide 7: Electromagnetic Spectrum
Laser wavelengths are usually in the Ultraviolet, Visible or Infrared Regions of the Electromagnetic Spectrum.
Notes: <number>

## Slide 8: Common Ultraviolet Lasers
Common Ultraviolet Lasers
Ultraviolet (UV) radiation ranges from 200-400 nm.
10
-12
10
-10
10
-8
10
-6
10
-4
10
-2
10
0
10
2
10
4
10
6
10
8
x-rays
gamma rays
Ultra-
violet
Infrared
Radar
Radio
waves
Electric
waves
Ionizing
Radiation
Wavelength (cm)
Common Ultraviolet Lasers |  |  |  |  |  | 
Argon fluoride | Krypton chloride | Krypton fluoride | Xenon chloride | Helium cadmium | Nitrogen | Xenon fluoride
193 nm | 222 nm | 248 nm | 308 nm | 325 nm | 337 nm | 351 nm
Ultra-
violet
Ultra-
violet

## Slide 9: Common Infrared Lasers
10
-12
10
-10
10
-8
10
-6
10
-4
10
-2
10
0
10
2
10
4
10
6
10
8
x-rays
gamma rays
Ultra-
violet
Infrared
Radar
Radio
waves
Electric
waves
Ionizing
Radiation
Wavelength (cm)
Infrared radiation ranges from 760-10,000 nm.
Infrared
Infrared
Common Infrared Lasers |  |  |  |  |  |  |  | 
Near Infrared |  |  |  |  |  |  | Far Infrared | 
Ti Sapphire | Helium neon | Nd:YAG | Helium neon | Erbium | Hydrogen fluoride | Helium neon | Carbon dioxide | Carbon dioxide
800 nm | 840 nm | 1,064 nm | 1,150 nm | 1,504 nm | 2,700 nm | 3,390 nm | 9,600 nm | 10,600 nm

## Slide 10: Common Visible Light Lasers
Violet | Helium cadmium | 441 nm
Blue | Krypton | 476 nm
 | Argon | 488 nm
Green | Copper vapor | 510 nm
 | Argon | 514 nm
 | Krypton | 528 nm
 | Frequency doubled Nd YAG | 532 nm
 | Helium neon | 543 nm
Yellow | Krypton | 568 nm
 | Copper vapor | 570 nm
 | Rohodamine 6G dye (tunable) | 570 nm
 | Helium neon | 594 nm
Orange | Helium neon | 610 nm
Red | Gold vapor | 627 nm
 | Helium neon | 633 nm
 | Krypton | 647 nm
 | Rohodamine 6G dye | 650 nm
 | Ruby (CrAlO3) | 694 nm
The wavelength range for light that is visible to the eye ranges from 400-760 nm.
Notes: <number>

## Slide 11: Biology of the Eye
Cornea
Lens
Fovea (focal point)
Retina
Notes: <number>

## Slide 12: Retinal Hazard Region
The wavelength range of light that can enter the eye is 400 to 1400 nm, though the range that we can actually see is only 400 – 760 nm.
The eye can focus a collimated beam of light to a spot 20 microns in diameter on the retina (called the focal point).
This focusing ability places the retina at risk when exposed to laser light in the wavelength range that will penetrate to the retina, because even fairly low wattage laser light can impact the retina with 100,000 times the radiant power that entered the eye. Because of this optical gain, laser light in the 400 – 1400 nm is referred to as the Retinal Hazard Region.
This is important to remember when working with infrared lasers, because the retina can be injured even though the laser is invisible.
Notes: <number>

## Slide 13: Retinal Hazard Region Eye Injury
Visible and Near-Infrared
Far-Infrared and Far-Ultraviolet
Near Ultraviolet
Notes: <number>

## Slide 14: Biological Hazards - Retina
Thermal damage to the retina occurs in the Retinal Hazard Region (from 400 nm – 1400 nm). Thermal damage is not cumulative, as long as the retina cools down between exposures.
Photochemical damage is severe at shorter visible wavelengths (violet & blue) and is cumulative over a working day.
Acoustic shock from exposure to high energy pulsed lasers results in physical tissue damage.
Notes: <number>

## Slide 15: Biological Hazards – Cornea & Lens
Inflammation injury to the cornea is caused by ultraviolet (UV) wavelengths (200-400 nm). This is the same type of injury that is caused by snow blindness.
Chronic exposure can cause cataract formation in the lens of the eye just as UV from the sun does.
Cornea
Lens
Notes: <number>

## Slide 16: Biological Hazards - Skin
Ultraviolet (UV)
  UV can cause skin injuries comparable to sun burn.
  As with damage from the sun, there is an increased risk for developing skin cancer from UV laser exposure.
Thermal Injuries
  High powered (Class 4) lasers, especially from the infrared (IR) and visible range of the spectrum, can burn the skin.
Notes: <number>

## Slide 17: Laser Safety Regulations
Occupational Safety & Health Administration (OSHA)
  No specific laser safety regulations, but will cite safety issues under the General Duty Clause and will enforce the ANSI standard for laser safety.
American National Standards Institute (ANSI)
  ANSI Z136.1 (2007) Safe Use of Lasers
  ANSI Z136.2-6 Specific references
Notes: <number>

## Slide 18: Laser Hazard Classes
The ANSI Laser Safety standard has defined Laser Hazard Classes, which are based on the relative dangers associated with using these lasers.
Class 1
Class 1M
Class 2
Class 2M
Class 3R
Class 3B
Class 4
Most Hazardous
Least Hazardous

## Slide 19: Class 4 Lasers
Visible and near-IR lasers will cause severe retinal injury and burn the skin. Even diffuse reflections can cause retinal injuries.
UV and far-IR lasers of this class can cause injury to the surface of the eye and the skin from the direct beam and specular reflections.
Even a diffuse reflection can cause injury.
May also produce laser generated air contaminants (LGAC).
This class of laser can cause fires.
Requires a Laser Safety Officer and written Standard Operating Procedures.
Notes: <number>

## Slide 20: Laser Safety Officer (LSO)
The Laser Safety Officer (LSO) is someone who has authority to monitor and enforce the control of laser hazards and effect the knowledgeable evaluation and control of laser hazards.
The LSO for GWU is Dan Hibbing with the Office of Laboratory Safety (OLS).
Designation of a Deputy LSO (DLSO) will vary according to the laser in question and usually be a faculty member or operator who has primary responsibility for routine work.
Notes: <number>

## Slide 21: Maximum Permissible Exposure (MPE)
The Maximum Permissible Exposure (MPE) is the highest level of radiation to which a person can be exposed without hazardous effects.
The MPE is specified in W/cm2 for continuous wave lasers and in J/cm2 for pulsed lasers. The value depends on wavelength, exposure duration and pulse repetition frequency.
Notes: <number>

## Slide 22: Non-Beam Hazards
Non-beam hazards refer to anything other than the laser itself that can create a hazard. This type of hazard includes:
  Electrical Hazards
  Fire Hazards
  Laser Generated Air Contaminants (LGAC)
  Compressed Gases
  Noise
Notes: <number>

## Slide 23: Control Measures
There are several measures that can be taken to prevent injury from lasers. These measures include:
  Engineering Controls
  Administrative Controls
  Personnel Protective Equipment
  Warning Signs and Labels
Notes: <number>

## Slide 24: Engineering Controls
Engineering controls are measures that are incorporated into the laser system and are designed to prevent injury to personnel. Engineered safety controls are preferable to PPE or Administrative controls.
Examples include
  Protective housings
  Interlocks on Removable protective housings
  Service access panels
  Key control master switch (Class 3B & 4)
  Viewing Windows, Display Screens, Collecting Optics
  Beam path enclosures
  Remote interlock connectors (Class 3B & 4)
  Beam Stop or attenuator (Class 3B & 4)
Notes: <number>

## Slide 25: Administrative Controls
Designation of Nominal Hazard Zones (NHZ).
Written Standard Operating Procedures (SOP’s) which are enforced by the LSO and DLSO.
Warning signs at entrances to room.
Training for all personnel who will be operating a laser or in the vicinity of the laser while it is in operation. (Training is also required for those using Class 2 and 3R lasers.)
Allow only authorized, trained personnel in the vicinity of the laser during operation.
Class 4 Laser
Standard Operating Procedures
Administrative controls are procedures that are designed to prevent personnel from injury. Examples of administrative controls required for Class 3B & 4 lasers include:

## Slide 26: PPE for Skin
Ultraviolet lasers and laser welding/cutting operations may require that tightly woven fabrics be worn to protect arms and hands.
For lasers with wavelengths > 1400 nm, large area exposures to the skin can result in dryness and even heat stress.
Personnel Protective Equipment (PPE) for Skin exposed to Class 3B or 4 lasers:
Notes: <number>

## Slide 27: PPE for Eyes
PPE is not required for class 2 or 3R lasers unless intentional direct viewing > 0.25 seconds is necessary.
Personnel Protective Equipment (PPE) for eyes exposed to Class 3B or 4 lasers is mandatory. Eyewear with side protection is best. Consider these factors when selecting eyewear:
  Optical Density (OD) of the eyewear
  Laser Power and/or pulse energy
  Laser Wavelength(s)
  Exposure time criteria
  Maximum Permissible Exposure (MPE)
  Filter characteristics
Notes: <number>

## Slide 28: Protect Your Eyes!
In a fraction of a second, your vision can go dark.
Notes: <number>

## Slide 29: Warning Labels
Class 2: “Laser Radiation – Do Not Stare into Beam.”
Class 2M: “Laser Radiation – Do not Stare into Beam or View Directly with Optical Instruments.”
Class 3R: “Laser Radiation – Avoid Direct Eye Exposure.”
Class 3B:
  “Laser Radiation – Avoid Direct Exposure to Beam”
Class 4: “Laser Radiation – Avoid Eye or Skin Exposure to Direct or Scattered Radiation.”
Only Class 1 lasers require no labels. All other lasers must be labeled at the beam’s point of origin.
Notes: <number>

## Slide 30: Warning Signs
Warn of the presence of a laser hazard in the area
Indicate specific laser safety policies
Indicate the relative hazard such as the Laser Class and the location of the Nominal Hazard Zone
Indicate precautions needed such as PPE requirements for eyewear, etc.
All rooms with class 3R, 3B or 4 lasers must have appropriate signs posted at all entrances. Signs must:
Notes: <number>

## Slide 31: Laser Warning Signs
“DANGER” indicates a very dangerous situation that could result in serious injury or death. This sign should be used for Class 3R, 3B, and 4 lasers.
“CAUTION” indicates a potentially hazardous situation which could cause a less serious injury. This sign should be used for Class 2 and 2M lasers.
“NOTICE” does not indicate a hazardous situation. This sign should only be used to make people aware of facility policies regarding laser safety and/or to indicate that a service operation is in progress.
Notes: <number>

## Slide 32: “CAUTION” Warning Sign
CAUTION
Laser Class and system
Type of Laser, emitted wavelength, pulse duration, and maximum output
Safety Instructions may include:
  Eyewear Required
  Invisible laser radiation
  Knock Before Entering
  Do Not Enter When Light is On
  Restricted Area
Safety Instructions
Notes: <number>

## Slide 33: “DANGER” Warning Sign
DANGER
Laser Class and system
Type of Laser, emitted wavelength, pulse duration, and maximum output
Safety Instructions may include:
  Eyewear Required
  Invisible laser radiation
  Knock Before Entering
  Do Not Enter When Light is On
  Restricted Area
Safety Instructions
Notes: <number>

## Slide 34: “NOTICE” Sign for Laser Repair
“NOTICE” Sign for Laser Repair
NOTICE
Safety Instructions (such as “Laser Repair in Progress”)
Laser Class and system
Type of Laser, emitted wavelength, pulse duration, and maximum output
Safety Instructions may include:
  Eyewear Required
  Invisible laser radiation
  Knock Before Entering
  Do Not Enter When Light is On
  Restricted Area
Notes: <number>

## Slide 35: Medical Surveillance
Medical surveillance may be available through the department for individuals who work with or around Class 3B and 4 lasers.
Contact the OLS for assistance in arranging a medical evaluation.
Notes: <number>

## Slide 36: For More Information
The ANSI Z136.1 Laser Safety Standard is the best reference to consult for laser safety information.
  One copy is available for viewing in Ross Hall 627, OLS (This copy will not be loaned out.)
  You can also purchase a copy of the standard through the Laser Institute of America web site (http://www.laserinstitute.org).
Call OLS at 4-2630 if you have additional questions about laser safety.
Notes: <number>